## Slide 1: vt01 _ sqb_8wire
data from 20200121

## Slide 2
Data taken from this die

## Slide 3
measured this device

## Slide 4
Here’s the zoom

## Slide 5
Even closer view

## Slide 6
I-V curve of the SQUID
Ic close to design
Hysteretic due to beta_c being > 1 resistance higher than expected

## Slide 7
Zoom of the hysteretic region

## Slide 8
Flux modulation for many values of SQUID current bias

## Slide 9
Flux modulation for best bias; linear trend due to non-ideal 4-wire

## Slide 10
After correcting for the linear trend, extract modulation amplitude and mutual inductance.
Modulation amplitude should be closer to 50 uV; again I suspect beta_c playing a role.
Mutual inductance designed to be 6.3 pH; acceptable error, probably due to layer thicknesses.